## GEX Form 100-269(a) Rev. B
| Test Method: | 100-270
| Instrument S/N: |  |  | Photometric Performance at 465nm (values are in absorbance units)
| Instrument Model: | GENESYS 30 |  | Standard ID | Certified Value | Uncertainty | G30 Spec | Lower Limit | Upper Limit | Value As Found | Pass/Fail
| Date: |  |  | SA0799-1 | 0.2889 | 0.0026 | 0.002 | 0.2843 | 0.2935 |  | FAIL | 22
| Relative Humidty (%): |  |  | SA0799-2 | 0.4765 | 0.0026 | 0.0024 | 0.4715 | 0.4815 |  | FAIL
| Temperature (°C): |  |  | SA0799-3 | 1.0334 | 0.0031 | 0.0052 | 1.0251 | 1.0417 |  | FAIL
| Calibration Standards Set ID#: |  |  | SA0799-4 | 1.5356 | 0.007 | 0.0077 | 1.5209 | 1.5503 |  | FAIL
| Calibration Certificate #: |  |  | Photometric Performance at 590nm (values are in absorbance units)
Notes: |  |  |  | Standard ID | Certified Value | Uncertainty | G30 Spec | Lower Limit | Upper Limit | Value As Found | Pass/Fail
|  |  |  | SA0799-1 | 0.3004 | 0.0026 | 0.002 | 0.2958 | 0.305 |  | FAIL
|  |  |  | SA0799-2 | 0.5017 | 0.0026 | 0.0025 | 0.4966 | 0.5068 |  | FAIL
|  |  |  | SA0799-3 | 1.0968 | 0.0031 | 0.0055 | 1.0882 | 1.1054 |  | FAIL
|  |  |  | SA0799-4 | 1.6306 | 0.007 | 0.0082 | 1.6154 | 1.6458 |  | FAIL
|  | Wavelength Accuracy Test at 525nm (nominal) |  |  |  |  | Wavelength Accuracy Test at 782nm (nominal)
[5 empty] | Measured Data (A) | (nm) |  |  |  | Measured Data (A) | (nm)
[6 empty] | 515 |  |  |  |  | 770
[6 empty] | 516 |  |  |  |  | 771
[6 empty] | 517 |  |  |  |  | 772
|  |  | Standard ID: | SA0799-5 |  | 518 |  | Standard ID: | SA0799-5 |  | 773
|  |  | Certified Value (nm): | 521.8 |  | 519 |  | Certified Value (nm): | 777.2 |  | 774
|  |  | Uncertainty (nm): | 1 |  | 520 |  | Uncertainty (nm): | 1 |  | 775
|  |  | G30 Spec (nm) | 2 |  | 521 |  | G30 Spec (nm) | 2 |  | 776
|  |  | Lower Limit (nm): | 519 |  | 522 |  | Lower Limit (nm): | 774 |  | 777
|  |  | Upper Limit (nm): | 525 |  | 523 |  | Upper Limit (nm): | 780 |  | 778
[6 empty] | 524 |  |  |  |  | 779
|  |  | Peak As Found (nm): | #N/A |  | 525 |  | Peak As Found (nm): | #N/A |  | 780
|  |  | Pass/Fail: | #N/A |  | 526 |  | Pass/Fail: | #N/A |  | 781
[6 empty] | 527 |  |  |  |  | 782
[6 empty] | 528 |  |  |  |  | 783
[6 empty] | 529 |  |  |  |  | 784
| Performed By: |  |  |  |  | Date:
ENTER NAME - Q.A. Technician
| Reviewed By: |  |  |  |  | Date:
ENTER NAME - Q.A. Manager

## GEX Form 100-269(b) Rev. B
| Test Method: | 100-271 Method 2
| Instrument S/N: |  | Photometric Performance at 465nm (values are in absorbance units)
| Instrument Model: | Evolution 220 | Standard ID | Certified Value | Uncertainty | Evo220 Spec | Lower Limit | Upper Limit | Value As Found | Pass/Fail
| Date: |  | SA0799-1 | 0.2889 | 0.0026 | 0.006 | 0.2803 | 0.2975 |  | FAIL | 22
| Relative Humidty (%): |  | SA0799-2 | 0.4765 | 0.0026 | 0.006 | 0.4679 | 0.4851 |  | FAIL
| Temperature (°C): |  | SA0799-3 | 1.0334 | 0.0031 | 0.006 | 1.0243 | 1.0425 |  | FAIL
| Calibration Standards Set ID#: |  | SA0799-4 | 1.5356 | 0.007 | 0.006 | 1.5226 | 1.5486 |  | FAIL
| Calibration Certificate #: |  | Photometric Performance at 590nm (values are in absorbance units)
Notes: |  |  | Standard ID | Certified Value | Uncertainty | Evo220 Spec | Lower Limit | Upper Limit | Value As Found | Pass/Fail
|  |  | SA0799-1 | 0.3004 | 0.0026 | 0.006 | 0.2918 | 0.309 |  | FAIL
|  |  | SA0799-2 | 0.5017 | 0.0026 | 0.006 | 0.4931 | 0.5103 |  | FAIL
|  |  | SA0799-3 | 1.0968 | 0.0031 | 0.006 | 1.0877 | 1.1059 |  | FAIL
|  |  | SA0799-4 | 1.6306 | 0.007 | 0.006 | 1.6176 | 1.6436 |  | FAIL
| Performed By: |  |  |  | Date:
ENTER NAME - Q.A. Technician
| Reviewed By: |  |  |  | Date:
ENTER NAME - Q.A. Manager

## GEX Form 100-269(c) Rev. B
| Test Method: | 100-271 Method 3
| Instrument S/N: |  | Photometric Performance at 465nm (values are in absorbance units)
| Instrument Model: | Evolution 220 | Standard ID | Certified Value | Uncertainty | Evo220 Spec | Lower Limit | Upper Limit | Value As Found | Pass/Fail
| Date: |  | SA0799-1 | 0.2889 | 0.0026 | 0.006 | 0.2803 | 0.2975 |  | FAIL | 22
| Relative Humidty (%): |  | SA0799-2 | 0.4765 | 0.0026 | 0.006 | 0.4679 | 0.4851 |  | FAIL
| Temperature (°C): |  | SA0799-3 | 1.0334 | 0.0031 | 0.006 | 1.0243 | 1.0425 |  | FAIL
| Calibration Standards Set ID#: |  | SA0799-4 | 1.5356 | 0.007 | 0.006 | 1.5226 | 1.5486 |  | FAIL
| Calibration Certificate #: |  | Photometric Performance at 590nm (values are in absorbance units)
Notes: |  |  | Standard ID | Certified Value | Uncertainty | Evo220 Spec | Lower Limit | Upper Limit | Value As Found | Pass/Fail
|  |  | SA0799-1 | 0.3004 | 0.0026 | 0.006 | 0.2918 | 0.309 |  | FAIL
|  |  | SA0799-2 | 0.5017 | 0.0026 | 0.006 | 0.4931 | 0.5103 |  | FAIL
|  |  | SA0799-3 | 1.0968 | 0.0031 | 0.006 | 1.0877 | 1.1059 |  | FAIL
|  |  | SA0799-4 | 1.6306 | 0.007 | 0.006 | 1.6176 | 1.6436 |  | FAIL
|  |  | Wavelength Accuracy Test at 525nm (nominal) |  |  |  | Wavelength Accuracy Test at 782nm (nominal)
|  |  |  | Standard ID: | SA0799-5 |  |  | Standard ID: | SA0799-5
|  |  |  | Certified Value (nm): | 521.8 |  |  | Certified Value (nm): | 777.2
|  |  |  | Uncertainty (nm): | 1 |  |  | Uncertainty (nm): | 1
|  |  |  | Evo220 Spec (nm) | 0.8 |  |  | Evo220 Spec (nm) | 0.8
|  |  |  | Lower Limit (nm): | 520.2 |  |  | Lower Limit (nm): | 775.2
|  |  |  | Upper Limit (nm): | 523.8 |  |  | Upper Limit (nm): | 778.8
|  |  |  | Peak As Found (nm): |  |  |  | Peak As Found (nm):
|  |  |  | Pass/Fail: | FAIL |  |  | Pass/Fail: | FAIL
| Performed By: |  |  |  | Date:
ENTER NAME - Q.A. Technician
| Reviewed By: |  |  |  | Date:
ENTER NAME - Q.A. Manager